This guide specification was prepared utilizing 3-part format recommended by the Construction Specifications Institute (CSI), and generally incorporates recommendations from their SectionFormat™/Page Format™, and MasterFormat™, latest Editions, insofar as practicable. Carefully review and edit the text to meet the Project requirements and coordinate this Section with the remainder of the Specifications and the Drawings. Where bracketed text is indicated, e.g. [text], make appropriate selection and delete the remainder of text within additional brackets, highlighting, and bold face type, if any. Consult the manufacturer for assistance in editing this guide specification for specific Project applications where necessary. This Specification was current at the time of publication but is subject to change. Please confirm the accuracy of these specifications with the manufacturer prior to use.

Section: 11 53 13

Vertical Laminar Air Flow (VLAF)

Type ha polypropylene fume hood (high Acid Use)

PART 1 - GENERAL

SUMMARY

Section Includes:

Benchtop polypropylene vertical laminar flow total exhaust fume hood.

SPECIFIER: Revise subparagraphs below to suit Project requirements.

Related accessories

Junction Boxes, Switches, Receptacles, and Other Controls

Fixture, Sinks or Accessories

Work Surfaces

Fume hood base

Filler panels and ceiling enclosures

Demonstration and training in the use and maintenance of the fume hoods.

Related Requirements:

- SPECIFIER: Revise section numbers and titles in subparagraphs below per CSI MasterFormat and Project requirements.

Drawings and general provisions of the Contract, including General and Supplementary Conditions and Division 11 Specification Sections, apply to this Section

- SPECIFIER: Delete Section 01 21 00 - Allowances if locations are clearly shown on the Drawings and sizes are either indicated alongside the locations or specified herein.

Section 01 21 00 – Allowances: For cash or quantity allowances covering access panels and frames.

Division 01 33 00 – Submittal Procedures: For administrative and procedural requirements for processing of submittals during the construction phase.

Division 01 77 00 – Closeout Procedures: For administrative and procedural requirements for completion of the Work.

SPECIFIER: Retain subparagraph below to suit Project requirements; revise section number and title per project requirements.

Division 12 35 00 Manufactured Plastic Casework

Division 22 Plumbing: Furnishing and installation of plumbing utilities and final connections.

Division 23 Heating, Ventilating and Air Conditioning: Furnishing and installation of exhaust duct work and equipment and final connections to fume hoods.

Division 26 Electrical: Furnishing and installation of electrical utilities and final connections.

References

- American National Standards Institute/American Industrial Hygiene Association (ANSI/AIHA)

ANSI/AIHA Z9.5 – 2022, Laboratory Ventilation

American National Standards Institute/American Society of Heating, Refrigeration, and Air Conditioning Engineers (ANSI/ASHRAE)

ANSI/ASHRAE Standard 110 – 2016, Method of Testing Performance of Laboratory Fume Hoods

American National Standards Institute/International Organization for Standardization/American Society for Quality Control (ANSI/ISO/ASQC)

ANSI/ISO/ASQC Q9001:2015

ASTM International

ASTM C1048-04, Standard Specification for Heat-Treated Glass—Kind HS, Kind FT Coated and Uncoated Glass

ASTM C1172-24, Standard for Laminated Architectural Flat Glass

ASTM D570-22 (2010) e1, Standard Test Methods for Water Absorption of Plastics

ASTM D638-14, Standard Test Methods for Tensile Properties of Plastics

ASTM D695-23, Standard Test Methods for Compressive Properties of Rigid Plastics

ASTM D790-17, Standard Testing Methods for Flexural Properties of Unreinforced and Reinforced Plastics and Electrical Insulating Materials

ASTM D4101-24, Stand Specification for Polypropylene Injection and Extrusion Materials

ASTM E162-25a, Standard Method for Flammability of Materials Using a Radiant Heat Source

The Scientific Equipment and Furniture Association (SEFA)

SEFA 1, Recommended Practices for Laboratory Fume Hoods

SEFA 2, Recommended Practices for Laboratory Installation

SEFA 3, Recommended Practices for Laboratory Works Surfaces

SEFA 7, Recommended Practices for Fixtures

SEFA 8PL, Recommended Practices for Laboratory Grade Polypropylene Casework

SEFA 10, Recommended Practices for Laboratory Adaptable Casework Systems

SEFA 11, Recommended Practices for Liquid Chemical Storage

- SPECIFIER:  Delete subparagraph below if no fire-rated fume hoods are specified.

Underwriters Laboratories, Inc. (UL)

UL 1805, Standards for Laboratory Hood and Cabinets

- US Consumer Product Safety Commission (CPSC) THIS IS OBSOLETE – USE IBC?

16 CFR Part 1201 (1-1-05 Edition) Safety Standards for Architectural Glazing Materials

Canadian Standards Associations (CAN/CSA)

CANS/CSA C22.2 No. 61010-1 Safety Requirements for Electrical Equipment for Measurement, Control and Laboratory Use

Action InformatIon and Submittals

- Submit in accordance with Division 11

Product Data: For each product indicated, submit technical data. Including the following:

Manufacturer model number.

Detailed specifications of construction.

Accessories and components that will be included for Project.

Shop Drawings: Include plans, elevations, and sections with dimensions, description of materials and finishes, general construction rough-in dimensions, component connections, anchorage methods, hardware, utility service requirements, and attachments to other work.

Include layout of units with relation to surrounding walls, doors, windows, lighting and air conditioning fixtures, connections of hood-to-hood exhaust system, location of access doors, and junction boxes.

Indicate clearance requirements for access and maintenance.

Indicate utility service connections for water, drainage, and power: include roughing-in dimensions.

Coordinate shop drawings with other trades whose work affects installation or performance of fume hoods.

Operations and Maintenance Manuals.

Samples:

Not less than 4” square piece of polypropylene stock used in the general construction of the fume hood, in thickness and finish specified, if requested by the architect.

Not less than 4” square piece of work surface material, in thickness, color, and finish specified.

Test and Evaluation Reports: Submit results of testing required under articles 2.09 SOURCE QUALITY CONTROL and 3.03 FIELD QUALITY CONTROL

Qualification Statements: Written statement from fume hood manufacturer, the installer meets requirements specified in Article 1.06 Quality Assurance.

Closeout submittals

- Submit in accordance with Section

Operation and Maintenance Data: Include detailed instructions for use and maintenance of fume hoods.

Warranty: Submit (2) copies for inclusion in the Owner’s systems and equipment maintenance manual.

Supply cleaning information for different materials used.

Supply part numbers of replaceable equipment, such as handles, hinges and screws.

Quality Assurance

- Compliance:

Comply with the provisions of the Building Code, these specifications, and standards referenced in Article 1.03 REFERENCES, except where more stringent requirements are shown on the Drawings or specified herein.

Fume hoods and their installations shall comply with laws, ordinances, rules, regulations and orders of governing authorities having jurisdiction over this portion of the Work.

Manufacturer: A firm with undivided responsibility for the fabrication of fume hoods, performance at a single location, and with minimal exposure to outside contaminants.

Installer: The manufacturer or an installer authorized by the manufacturer to install laboratory equipment with at least 10 years of successful experience in installing fume hoods similar to those specified.

Single Source Responsibility: Fume Hoods and accessories included in this section shall be manufactured or furnished by the same manufacturer or laboratory furniture supplier for single responsibility.

Delivery, storage, and handling

- Deliver, store and handle fume hoods using means and methods that will prevent damage, deterioration, and loss, including theft and vandalism. Comply with manufacturer’s written instructions.

Deliver fume hoods to Project site in undamaged condition in manufacturer’s original sealed containers or other packaging system, complete with labels and instructions for handling, storing, unpacking, protecting, and installing.

Inspect fume hoods on delivery to determine compliance with the Contract Documents and to determine that products are undamaged and properly protected.

The Construction Manager will provide a secure location and enclosure at Project site for storage of fume hoods until time of installation.

Field Conditions

- Environmental Requirements:

Do not install fume hoods until building is enclosed, and surrounding construction is completed including overhead work associated with the fume hoods, floor, wall and ceiling finishes, mechanical, electrical, plumbing and fire protection work performed by others.

Warranty

- Warranty Period: One year from the date of customer acceptance or substantial completion, whichever is later.  Stipulate that defects that develop within the Warranty period shall be removed, repaired, or replaced at no additional cost to owner.

Failures include but are not limited to:

Manufacturer defects

Structural failure

Warping

Finish

PART 2 - PRODUCTS

Fume Hood

- Acceptable Manufacturers:

- LabAire Systems (LAS), a division of Activar Plastics Products Group

- 9650 Newton Avenue South

- Bloomington, MN 55431

- 877-410-4910

- www.labairesystems.com

- info@labairesystems.com

- Substitutions: Manufacturers seeking approval for their products are required to comply with the Owner’s Instructions to Bidders, generally contained in the Project Manual.

Manufactured units

- Fume Hood: Model NU-156 FumeGard Vertical laminar flow total exhaust polypropylene hood based on design previously manufactured by NuAire, Inc.

- External Hood Widths:

- SPECIFIER:  Delete Model Not Required

4-foot Cabinet, 690CFM (24” Depth) @ 0.7” W.G or 790 CFM (30” Depth)

- @ 0.7” W.G.

5-foot Cabinet, 910CFM (24” Depth) @ 0.8” W.G. or 1040CFM (30” Depth)

- @ 0.9” W.G.

6-foot Cabinet, 1120CFM (24” Depth) @ 1.0” W.G. or 1280CFM (30” Depth

- @1.1” W.G.

8-foot Cabinet, 1550CFM (24”Depth) @ 1.2” W.G. or 1770CFM (30” Depth)

- @1.3”W.G.

Work Surface

Depth:

24” Inside Work Surface Depth with an External Hood Depth 36”

30” Inside Work Surface Depth with an External Hood Depth 42”

Performance Criteria

- General Performance Requirements: Properly installed and operating cabinets serve as ventilated workspaces, with the purpose of capturing, containing, and evacuating fumes, vapors and other airborne particulates associated with the activity within the enclosure. For a design point the Fume Hood has a minimum average inflow velocity through the work access opening at 105 FPM with a 10-inch nominal sash opening. Working sash heights shall vary between 8-12 inches without loss of “personnel protection factor”. The maximum non-working height shall be 20-inches.

- SPECIFIER:  Non-working heights up to 30-inches may be specified at additional cost.

- Illumination within the Work Area:

Definition of Work Area: The area from the work surface to the maximum opening height of the front sash to the front face of the baffles.

The average illumination level within the work area shall be 90 to 120-foot candles

Air Flow: Proper Design shall allow for safe and consistent air flow inside the hood and around the hood face.

Cabinet is designed as a bench style, single pass flow-through design in which all HEPA filtered work zone (downflow) and work access inflow air, is drawn through the cabinet's internal exhaust plenums to a single exit point at the top.

The system shall maintain a safe face velocity based on the recommendations of the manufacturer.

Air shall be drawn into the HEPA filter by the internal supply blower from a prefilter at the top of the cabinet.  The air shall exit the HEPA filter through a diffuser into the work zone in a laminar manner. The air shall split at the removable work surface with ½ being drawn into a rear inlet grill and ½ drawn into the front inlet grill along with the work access inflow.  The combination of the downflow and inflow into the front inlet grill forms an “air barrier” that creates the “personnel protection factor.

The supply HEPA filter shall be the full cabinets’ work surface width and depth and removable from the front of the cabinet. The work zone shall be the same fixed cross-sectional area as the HEPA filter.

The supply HEPA filter shall be epoxy coated, aluminum frame and constructed from non-woven pleated fiberglass media with separatorless media. All media shall be a minimum of 99.99% efficient on 0.3 micron particles.

- SPECIFIER: Revise subparagraphs below to suit Project requirements. Select specific option listed below at additional cost.

ULPA Filter 99.999% on 0.12 microns particles

Non-Boron, woven PTFE Media

The positive pressure supply plenum shall be surrounded by negative pressure, even down to the interface of the supply HEPA filter and the filter frame to absolutely preclude any gasket leaks from contaminating the HEPA filtered work zone air.

The downflow air velocity from a clean HEPA filter shall average from 55 to 65 LFPM with no single point more than 20% above or below the average velocity measured in a horizontal plane defined by the bottom of edge of the sash at nominal working height (10-inches).

The area beneath the removable work surface shall form a leak-tight exhaust plenum that shall drain to a single, polypropylene, 3/8-inch, ball-type drain valve. The drain valve shall be plumbed to a polypropylene “p-trap” where a sink/cup sink has been specified or dead ended into a ball valve.  1½- inch vent holes shall be provided through the bottom of the plenum/cabinet to provide venting for the base cabinets below.  One vent hole shall be provided for each base cabinet door and they shall be protected from plenum spills by 3-inch high “mushroom capped” stand pipes.

The internal downflow blower shall be a non-overloading, backward inclined, dynamically balanced, motorized impeller with the impeller and all hardware (i.e. nuts, bolts, screws, etc) TFE coated by a licensed coating manufacturer, for maximum corrosion resistance.  The coating shall be a minimum thickness of 1.5 mil.

All motors shall be permanent split capacitor type and sized to automatically compensate for airflow as the filter loads with particulate to achieve a fan delivery fall off of no more than 10% for an 80% increase in pressure drop across the filter. The Class B insulated motor shall be thermally protected, rated for 24-hour continuous operation, lubricated for life and speed controlled by a solid state, voltage regulated solid-state speed controller.

The exhaust volumes and negative static pressures required per cabinet size shall not exceed the following based on an average inflow velocity of 105 FPM at a work access opening height of 10-inches and an average internal downflow velocity of 65 LFPM

Vibration, due to the motor, shall be minimized through the use of vibration isolators when mounting the motor/blower assembly to the internal frame.

All electronic controls shall be housed in a separate gastight polypropylene enclosure and shall be accessible from the top of the cabinet.

Variations in air flow shall not exceed manufacturer’s in-house test results when tested in accordance with ANSI/ASHRAE 110.

Static Air Pressure Loss:

The design of the fume hood shall allow for minimal static pressure loss.

The maximum average static pressure loss reading shall be taken 3 diameters above the fume hood outlet, from 4 points, 90 degrees apart, and shall not exceed the guidelines in full, open position.

Materials

- Hood: Refer to manufacturer’s chart for physical characteristics.

Standard Polypropylene.

- SPECIFIER:  Select specific option listed below at additional cost

FM 4910 Flame Retardant Polypropylene

UL 94V0 Flame Retardant Polypropylene

PVDF

Halar

Work Surface:

Polypropylene

- SPECIFIER:  Select specific option listed below at additional cost

FM4910 Polypropylene

PVDF

Halar

Sash Material:

Clear AR (Abrasion Resistant)Polycarbonate (Tuffak) 6mm (1/4” Nominal)

- SPECIFIER:  Select specific option listed below at additional cost

Clear PVC (Polyvinylchloride) 6mm (1/4” Nominal)

Sash Counterweight Connection: Kevlar-reinforced polyurethane belt.

Adhesives/RTV/Sealants: 100 percent silicone.

Fasteners: Manufacturer’s standard recommended fasteners.

Polypropylene screws.

PTFE coated Stainless Steel screws (used outside work zone)

Steel shoulder bolts (used outside work zone, in chase area).

PVC Adhesive Caps (used for fastener covers)

Lighting

LED lights

UL Listed junction box and wiring.

Lights and Blower switch with on-off positions

- SPECIFIER:  Select specific option listed below at additional cost

Controls

- Sash Control

Manual raising and lowering.

Manual sash stop at 10”

Lighting

Electrical On/Off rocker switch.

- SPECIFIER:  Select specific option listed below at additional cost

Digital Manometer Monitor: Digital pressure sensor with LED Display shall monitor exhaust system negative or positive pressure. Pressure sensor shall be connected to the rear plenum, monitoring exhaust system negative pressure.

24 VAC

Operating pressure range +/- 0 o 1.99” w.g.

Operating pressure set point repeatability +/- 0.02” w.g.

Operating pressure accuracy +/-0.5% of reading

Operating temperature: 50◦F to 104◦F (10◦C to 40◦C)

Alarm

Audible, 73 dba

Visual, flashing yellow/red LED

Fixtures

- Sinks: Shall include strainer, tail piece and provide p-trap, plumbing to the rear of the hood base.

Molded Cup Sink

- SPECIFIER:  Select specific option listed below at additional cost

Molded Sink

Fabricated sink

Custom/ADA sink

Faucets:

Extruded Natural Polypropylene body material

No elastomer, metal and lubricant used in construction.

PTFE Seals.

Deck mount or wall mount available (subject to location).

Max flow rate of 2.5 GPM @ 80 PSI.

- Operating pressure range: 1-200 PSI.

SPECIFIER:  Select specific option listed below at additional cost

PVC material.

PVDF material.

Standard Vacuum or Atmospheric Vacuum Breaker.

Hot/Cold Mixing.

ADA Blade Handles.

Replaceable 0.2 Micron capsule filter.

Service Fixtures: Color-coded hose nozzle outlets mounted inside the fume hood and controlled from the exterior with color coded index handles for up to 8 fixtures.

Service lines shall be installed from valve to outlet.

Connections shall be made using Polypropylene compression fitting and Polyethylene tubing. Polypropylene compression fitting ½” MPT and 6ft of tubing will be provided for incoming services.

Valves

      - Air

      - Vacuum

      - Nitrogen

      - Deionized/Distilled water

      - Hot and/or cold water

Flammable gas shall be brass valve with aluminum piping, assembly to have CSA/ANSI Z21.15-CSA 9.1 listing.

Any external color-coded handles/rods shall be coated with a clear acid and solvent resistant plastic coating

Electrical services:

Three wire grounding type receptacles rated for 120-volt supply.

Outlet shall have a non-metallic, corrosive resistant, vapor-tight gasketed cover capable of being sealed in place with a gasketed thumbscrew and shall have removable inserts of plastic to seal outlet contacts when not in use.

Receptacles shall conform to NEMA standards.

- SPECIFIER: Select specific option listed below at additional cost

Three wire grounding type receptacles rated for 220-volt supply

Accessories

- SPECIFIER:  Select specific option listed below at additional cost

- Ceiling enclosure panels: Filler panels matching fume hood exterior to enclose space above fume hoods at front and sides of fume hoods and extending from tops of fume hoods to ceiling.

Panels shall be louvered (or slotted) to obtain air supply for the downflow, unless otherwise stipulated.

Side panels shall fasten to the top of the cabinet and rear wall. The front panel shall fasten only to the side panels and easily removable to provide access to the top for maintenance purposes.

Fasteners will be of PTFE coated stainless steel.

8ft ceiling height

10ft ceiling height

Custom height

PTFE DI/N2 spray Gun with 8ft of flexible tubing.

Deck Mount

Wall Mount

PTFE handheld eye wash gun with 8 ft of flexible tubing

Deck Mount

Wall Mount

Fume Hood Equipment Lattice Kit: All components are constructed of fiberglass and thread into receptacles in baffle.

Exhaust duct wash down.

Three, ¾-inch, schedule 80, polypropylene pipes (spray bars) extending horizontally across the rear exhaust duct, evenly spaced vertically and sealed on one end.

Spray bar shall consist of replaceable, 1/8-inch NPT, polypropylene spray heads, with a 110-degree spray angle and 0.28 GPM at 10 PSI, on evenly spaced centers (minimum of three), orientated down toward the floor.

Spray bars shall be manifolded to a common point through a ball valve mounted on the hood post as shown on the drawings for operator control (on/off and flow rate).

Valve shall be plumbed out the rear of the cabinet for connection to building utilities.  A 1½-inch, NPT threaded drain hole shall be installed in the center of the floor of the exhaust duct to remove washdown water. The floor shall ensure complete drainage of wash water.

Rear divider panel shall be removable for inspection and maintenance of the spray bar system.  The panel shall be gasketed via inlaid 1/8-inch-thick EPDM gasket material, around the entire periphery of the panel and fastened to the back via ¼-inch-20 flat head screws on 4-inch centers to form a leak-tight seal.

Exhaust damper with quadrant lock.

8” diameter

10” diameter

12” diameter

Fabrication

- General: The main assembly is constructed of ½-inch thick, white Polypropylene. Seams are fully seam-welded using hot air or nitrogen, forming a rigid internal structure, with removable, white Polypropylene access panels to enclose utilities and inner components. Panels are size differently to accommodate ease of use and function. Standard overall dimensional tolerances are plus-or-minus 1/8-inch. Where required each hood can be supplied with a Fume Hood base either in Polypropylene or Metallic Construction

- Joints: Joints are formed with 5/32-inch white Polypropylene welding rod.

Joints are watertight.

Exterior welds are shaved flush with construction material to create a uniform surface.

Structure: The upper structure of the cabinet containing the HEPA filter, blower, supply plenum and electronics shall use 1/2-inch thick polypropylene material and shall be an integral extension of the work space sides and back walls (i.e. one-piece construction).   Access panels shall be provided at the top of plumbing chase section for maintenance of sash pulleys, counterweight and Kelvar®-urethane belt.

The unit is constructed with the inner walls as the main support structure. Walls contain access ports to the utility areas, which are secured with ¼-20 white Polypropylene Philips head screws.  Ports are flush with the surrounding walls with lipped construction to reduce the potential of leakage.  External walls are removable and secured with ¼-20 PTFE coated Stainless Steel Philips head machine screws.

Access to the HEPA filters and motor blower assembly for maintenance and repair shall be through a flush mounted access panel in the upper portion of the hood face.  Access to internal components shall NOT require disassembly of the window sash and glide system.

Prefilter(s) shall be mounted on top of the motor/blower module and easily replaceable.  Prefilter(s) shall be 1-inch thick non-woven fiberglass and shall be 40 percent efficient per the Atmospheric Dust Test

Work Surface: The Hood work surface is dished for spill retention. The surface has a ribbed structure on the underside to add structural rigidity.

Viewing Panel and Sash: The front viewing panel is glazed with clear Polycarbonate. The frame supporting the sash is made of white Polypropylene and contains a ballast area. The sash will be attached with Kevlar-reinforced Polyurethane belt. The pulleys are made from Polypropylene. The handle is constructed of white Polypropylene.

The diffuser below the HEPA filter shall fit snugly full width and length of the work zone and constructed from 1/16-inch thick, 22 percent open, perforated PVC and held in place with polypro screws.  A vinyl wiper shall be installed between the sash and the interface of the diffuser and the upper workspace to provide a seal against the glazing material. The area above the wiper, but under the sash, shall be maintained under a negative pressure.

All internal electrical wiring shall be enclosed in flexible, UL Listed or Recognized, non-metal corrosive-resistant, liquid-tight conduit with Listed fittings.  All internal junction boxes and enclosures shall be constructed from polypropylene and shall be liquid-tight construction and gasketed where required. All exposed controls and visual indicators shall be constructed from non-metal corrosive-resistant materials or protected by non-metal, gas-tight enclosures with a clear viewing lens, where required.

The work surface shall be removable and constructed from ¼-inch thick polypropylene material, reinforced with polypropylene ribbing to maintain surface flatness.  The perforated front and rear sections shall be 3/16-inch diameter holes on 1/4-inch alternate centers, whose width is suitable to provide both personnel and product protection per the testing methods of NSF STD 49. Where sinks are called out, the work surface shall be segmented such that the sink section is fixed and separate from the rest of the removable work surface.

The work surface shall be removable and constructed from ¼-inch thick polypropylene material, reinforced with polypropylene ribbing to maintain surface flatness.  The perforated front and rear sections shall be 3/16-inch diameter holes on 1/4-inch alternate centers, whose width is suitable to provide both personnel and product protection per the testing methods of NSF STD 49. Where sinks are called out, the work surface shall be segmented such that the sink section is fixed and separate from the rest of the removable work surface.

Supply mains to the cabinet shall be connected to a UL Listed junction box of suitable size to accommodate required circuits as indicated on the hood schedule.  The junction box shall be non-metallic and require non-metallic liquid-tight connections.

Where required, all cabinets shall be provided with a separate support base constructed from ½-inch polypropylene. Fume hood bases shall be constructed per the requirements for polypropylene base cabinets, section 12350, except as specified herein

Fume hood bases shall have 1½-inch diameter adjustable levelers that are accessible from inside the base cabinet and protected from spills with a ½-inch high rim, covered with a removable cap.  A minimum of two (front and rear) levelers shall be provided for each 2-foot section of fume hood. The levelers shall thread through a minimum of 1½-inches of material for stability and permit an adjustment of 2-inches.

The top of the base cabinet shall have a 3-inch wide, full interior perimeter stretcher to facilitate assembly to the hood.

Hood bases shall have at least one sealed compartment where sinks/cup sinks are scheduled.  The remaining space shall contain a fixed shelf.  Where sinks are scheduled on both sides of the workspace, each shall have a sealed compartment. Each sealed compartment shall have an access panel for connection to building utilities.  Depending on hood size, any remaining space shall contain a fixed shelf.

Each segmented compartment shall be vented to the exhaust duct via a stand pipe, plumbed through the base of the VLAF cabinet to the exhaust duct.

All base cabinet doors shall have four 1½- inch diameter vent holes, centered 4-inches above the bottom of the door and evenly spaced.  An adjustable slide shall be located on the inside of the door, with an identical offset hole-pattern to control airflow through the compartment(s). For compartments with shelves, shelves shall permit air to pass through the front ½-inch of the shelf.

Source Quality Controls

Tests: Fume hoods are tested at the manufacturer’s assembly facility in compliance with ANSI/ASHRAE 110 (Manufacturer’s Test), by independent third-party testing prior to shipping; results are available to the Owner, if requested.

These tests generally include:

Tracer Gas Test

Face Velocity Test

Smoke Pattern Test

Saturation Smoke Test

PART 3 – Execution

Examination

- Examine substrates, areas, and conditions, with installer present where indicated for compliance with requirements for installation tolerances and other conditions affecting performance. Notify the Construction Manager of conditions detrimental to performance of the Work and recommended corrections. Where the installation and its completion will be delayed due to existing conditions, follow immediately with a written report. Proceed with installation only after unsatisfactory conditions have been resolved.

Installation

- General: Install fume hoods in strict accordance with manufacturer instructions, plumb, level, aligned, rigid, and securely anchored to supporting casework, in accordance with final shop drawings.

- Coordinate sequence of work with mechanical, plumbing, and electrical trades and with installation of related casework.

Field Quality Control

- Acceptance Testing: On completion of fume hood installation and before permitting regular use, perform acceptance tests in accordance with ANSI/ASHRAE 110, for face velocity and airflow smoke patterns.

Advise Owner and Architect in advance of dates and times that tests are to be performed on fume hoods.

Fume hoods are determined to have passed testing if test results meet performance criteria specified.

Fume hoods which fail to comply with scheduled tests shall be adjusted in the field by the appropriate mechanical trades and retested until compliance with specified limits of operation are obtained.

Cleaning and protection

- Clean fume hoods following installation in accordance with manufacturer instruction and recommendations using liquids that will not harm finishes and glazing.

Clean exposed surfaces of the fume hood including surfaces exposed to view, those within the work area, and glazing.

Advise Construction Manager of methods to protect fume hoods until acceptance by the Owner. Protect fume hoods from damage due to other construction activity, during installation and after acceptance.

Closeout activities

- Demonstration and Training: Engage a factory-authorized service representative to train Owner’s personnel to adjust, operate, and maintain fume hoods.

Perform training only after equipment has been installed, tested, and is operating correctly.

Job specific warranty information.

Supply cleaning information for different materials used.

Supply part numbers for replaceable equipment, such as handles, hinges and screws.

END OF SECTION 11 53 13